## Slide 1: Infant Respiratory Rate Monitor Group 11
Sean Ermer
Tim Lau
Leo Li
Mentor: Dr. Daniel Moran

## Slide 2: Need
Engineering World Health (EWH) Projects That Matter
Greatly improve mortality rates of infants worldwide
Not home monitoring
Apnea of Prematurity and SIDS
Cost efficient

## Slide 3: Need
4 million neonatal deaths per year
Chart: 
| Developed world | Developing world
Sales | 0.01 | 0.99

## Slide 4: Need
1.4 million deaths for Asphyxiation and SIDS
Chart: 
| Developed | Developing | Asphyxiation | SIDS
Sales | 0.01 | 0.76 | 0.23 | 0.12

## Slide 5: Specific Design Requirements
Attribute | Requirement
Size | Fit within a typical crib (1 m x 0.8 m x 0.5 m)
Mobility | Easily moved from crib to crib (Weigh < 3 lb)
Ease of Use | Does not require extensive mechanical/electrical knowledge
Recognizable and Audible | Able to draw nurse’s attention from reasonable distances
Accurate | Detect breathing cessation for > 20 sec
Reliable | Few false alarms
Non-invasiveness | Does not affect infant’s breathing
Cost | Must cost < $20 to manufacture and implement
Novelty | Does not infringe on existing patents

## Slide 6: Specific Design Requirements
Attribute | Requirement
Size | Fit within a typical crib (1 m x 0.8 m x 0.5 m)
Mobility | Easily moved from crib to crib (Weigh < 3 lb)
Ease of Use | Does not require extensive mechanical/electrical knowledge
Recognizable and Audible | Able to draw nurse’s attention from reasonable distances
Accurate | Detect breathing cessation for > 20 sec
Reliable | Few false alarms
Non-invasiveness | Does not affect infant’s breathing
Cost | Must cost < $20 to manufacture and implement
Novelty | Does not infringe on existing patents

## Slide 7: Design Components
Sensor
Interpreter
Alarm

## Slide 8: Design Components
Sensor
Interpreter
Alarm
Piezo film
Microcontroller
PCB mount buzzer

## Slide 9: Sensor
Piezoelectric model
  Materials that generate a voltage when deformed
  Relatively inexpensive (approx. $3/unit)
  Can generate relatively large currents

## Slide 10: Model

## Slide 11: Model
Monitor
Chest strap

## Slide 12: Chest Strap
71% polyester, 29% rubber
  Spring constant ~5 N/m
3cm width, variable circumference
Velcro for easy adjustment
Fabric “pocket” for piezo strip

## Slide 13: Monitor
Plastic casing
  Outside: LED, on/off button
  LED indicates low battery
Encloses circuitry for microcontroller, alarm
9V battery

## Slide 14: Product Components

## Slide 15: Circuit

## Slide 16: Circuit
Microcontroller
DC/DC converter
Voltage regulator

## Slide 17: Microcontroller
Metric | Weight | PIC10F220/222 | PIC12F510 | PIC16F506 | PIC12F615
Price | 10 | 9 | 9 | 7 | 5
RAM | 8 | 2 | 4 | 9 | 8
I/O Pins | 5 | 3 | 5 | 10 | 5
A/D Channels/Converters | 10 | 7 | 8 | 8 | 9
A/D Resolution | 7 | 8 | 8 | 8 | 10
Program Memory | 4 | 3 | 5 | 8 | 10
Total |  | 259 | 303 | 360 | 339

## Slide 18: Microcontroller
Metric | Weight | PIC10F220/222 | PIC12F510 | PIC16F506 | PIC12F615
Price | 10 | 9 | 9 | 7 | 5
RAM | 8 | 2 | 4 | 9 | 8
I/O Pins | 5 | 3 | 5 | 10 | 5
A/D Channels/Converters | 10 | 7 | 8 | 8 | 9
A/D Resolution | 7 | 8 | 8 | 8 | 10
Program Memory | 4 | 3 | 5 | 8 | 10
Total |  | 259 | 303 | 360 | 339

## Slide 19: Microcontroller: PIC16F506
Core functionality, inexpensive ($.62/unit)
14 pins
67 bytes RAM, 1.5kB program memory
High A/D resolution for comparator
4MHz oscillator
  Will be slowed down to 32kHz to save energy
  Gives ~38Hz sampling rate
  Infant breathing rate .5-1Hz

## Slide 20: Signal Acquisition
Piezo voltage corresponds to motion from breathing
  For max displacement motion, 4V/-4V
DC/DC converter changes -ve signal to +ve
  PIC16F506 cannot process -ve signal
Voltage regulator keeps steady input signal to microcontroller
  Stable 5V input
  Less opportunity for failure

## Slide 21: Alarm
Commercial 30mm PCB mount buzzer
Wide operating voltage range (1-30V)
Low current draw (~5mA)
95dB, 2.5kHz frequency sound

## Slide 22: Power Consumption
Microcontroller clocked at 32kHz
  38µA to run
Power regulator
  200µA
~5000 hrs (or 7 months) to drain battery if continuously on
  Assume energy cost of alarm, LED is negligible

## Slide 23: PCB
Cost: ~$7
3.1’’ x 2.8’’

## Slide 24: Software Flow Chart

## Slide 28: Safety
Most prominent safety issues:
  Noise above 80dB
  Power supply interruption
On/off switch and LED low battery indicator reduces these risks
Overall, very safe

## Slide 29: DesignSafe

## Slide 30: Assembly
Two options
Method 1: sell as complete package, no assembly required
  Outsource manufacturing
Method 2: sell all individual parts with instructions on how to assemble product
  Provides cheaper option for 3rd world countries

## Slide 31: Parts
Total: $9.61
PCB lead time: 2-3 weeks
Plastic casing lead time: 2-4 weeks
Weight: Monitor 2.5oz, chest strap ~1.7oz
Part | Price ($)
Microcontroller | 0.62
Resistors | 0.15
DC/DC converter | 1.11
Battery connector | 0.29
9V battery | 0.88
LED | 0.04
Part | Price ($)
Piezo strip | 2.95
Button | 0.18
Alarm | 0.89
Breadboard | 1.16
Wiring (20 gauge) | 0.04
Plastic casing | 1.2
Elastic strap | 0.10

## Slide 32: Conclusions
Met design specifications
Move on to working prototype, mass production
No IP
  EWH non-profit
  Disclosing design to public = lowest price
What did we learn?
What would we do different?

## Slide 33: Thank you!
Questions?